Infrared Characterization

System #1: DRIFTS

Manual

Francisco Zaera Group

prepared in January 2014,

and updated in September

by Ilkeun Lee

Table of Contents

1.	General Considerations/Overview of Equipment	4

2.	FT-IR Instrument and General Operation	4

a.	Spectrometer General Description	4

b.	FTIR Performance Characterization	4

i.	Measuring Spectral Signal-to-Noise	5

ii.	Maximizing Absorbance Signal	5

c.	Main FTIR Parameters	5

d.	Detector	6

i.	Types	6

ii.	Substitution	6

iii.	Preparation	7

iv.	Maintenance	7

e.	Optical Alignment	8

f. 	Transmittance spectra	9

i.	Choice of Background Spectra	9

g.	OPUS Software	9

i.	Spectra Acquisition	9

ii.	Data Processing	11

h. 	Maintenance	11

i.	IR Source	11

ii.	HeNe Laser	14

iii.	Gas Purging	16

3.	Gas Handling System	17

a. 	Design	17

b. 	General Operation Procedure	18

c. 	Gas and Liquid Sample Handling	18

d. 	Maintenance	18

e. 	Valves	19

f. 	Pressure Gauges	19

g. 	Mechanical Pumps	20

4.	Praying MentisTM DRIFTS Cell	21

a.	General Description	21

b.	Installation	21

i.	Open Beam Spectrum	21

ii.	Purge Installation	22

c.	Alignment	22

i.	Alignment	22

ii.	Verifying Throughput	23

d.	Low-Temperature Reaction Chamber (CHC) Installation	24

i.	Introduction	24

ii.	Observation Window Installation	25

iii.	Dome Installation	25

iv.	CHC Installation	26

v.	Fluid / Electrical Connections	28

e.	Temperature Controller (ATC) Connection	29

i.	Introduction	29

ii.	Connection	29

iii.	Setting Thermocouple Type	30

iv.	Setting Control Type	30

v.	Setting Units	30

vi.	Setting the Low Temperature Limit	30

vii.	Setting the High Temperature Limit	31

viii.	Setting the Low Power Limit	31

ix.	Setting the High Power Limit	31

x.	Setting the Auto-Tune Set Point	31

xi.	Auto-Tuning	32

f.	Sample Loading	32

i.	Collecting Background Spectrum	32

ii.	Sample Spectrum	32

g.	Maintenance	33

i.	Leak Checking	33

ii.	Replacing the O-ring	33

iii.	Cleaning the Reaction Chamber	33

iv.	Replacing the Heater	33

v.	Replacing the Thermocouple	34

5.	Typical Experiment Sequence	34

a.	Initial Steps	34

b.	Sample Preparation	34

c.	Background Spectrum Acquisition	34

d.	Gas Manifold Control	35

e.	Sample Spectrum Acquisition	35

f.	Temperature Control	36

g.	Final Steps	36

6.	Suggested Training for Beginner	36

7.	Reference Materials & Contacts	37

a.	Reference Materials	37

b. 	Contacts	37

# 1.	General Considerations/Overview of Equipment

Please read this manual carefully before using the DRIFTS system and keep it in a suitable place for future reference. Always follow the instruction described in this manual to ensure safety and to avoid damage. Improper use or failure to do safety instructions can result in serious injuries and/or property damage. Especially ensure that the CO alarm on the ceiling is ON when you use CO gas. If it doesn’t work, replace the battery inside or have another one prior to work.

DRIFTS system is placed on the bench in Chemical Science building room 137, and consists of six parts: FTIR spectrometer (Tensor 27, Bruker), Gas manifold with vacuum, DRIFT cell (Praying MentisTM, Harrick) Low-T reaction chamber (CHC, Harrick), Temperature controller (ATC, Harrick), and Computer (OPUS program).

Tensor 27 FTIR spectrometer is equipped with Mid-IR source, KBr beamsplitter, and DLaTGS detector. The DRIFT cell sits in the sample compartment and is purged with nitrogen gas. The CHC reactor is designed for operation from -150 oC (LN2 cooling) up to 600 oC (water cooling) under vacuum. The pressure range covers 10-6 to 103 Torr only, so do not try higher pressure to break the windows of the CHC dome. The ATC is operated by a process controller (Series 989, Watlow), so please refer its manual (0600-0009-0004 Rev B) also for the details. OPUS (ver. 4.2) program is installed in Windows XP professor (2002 version, service pack 3) computer. All the manuals should be stored in the manual cabinet, right next to Stan’s desk in the room 139, so please return them after you read.

# 2.	FT-IR Instrument and General Operation

## a.	Spectrometer General Description

The Fourier-transform infrared (FTIR) instrument used for this system is a Bruker Tensor 27 in Chemical Science building room 137.  More details about this instrument can be found in its manual (7th updated edition, September 2011).  General FTIR principles are well described in several books.  See, for instance: Peter R. Griffiths and James A.  de Haseth, "Fourier Transform Infrared Spectrometry", John Wiley & Sons, New York, 1986.

The DRIFTS system is placed in the sample compartment, with insulating tubes placed between the IR beam windows and the cell for nitrogen gas purging.

## b.	FTIR Performance Characterization

The FTIR spectrometer should be setup for optimal performance by choosing the appropriate parameters and aligning the sample and optics.

### i.	Measuring Spectral Signal-to-Noise

Signal-to-noise (S/N) ratio: This is a critical value dependent on the other parameters that should be minimized before performing experiments. It is checked by acquiring back-to-back spectra under identical conditions and ratioing those. S/N ratios can be calculated by the OPUS software (one of “Evaluate” menus), and should be done for two frequency regions, typically 2000–2200 cm-1 (the region with the least noise), and a second in a region of interest, around 3000 cm-1, for example.

### ii.	Maximizing Absorbance Signal

This is done by choosing a particular IR absorption peak in the spectra in the sample and following that by taking spectra as parameters are optimized.

## c.	Main FTIR Parameters

- Total intensity (Amplitude): This is measured by the peak-to-peak voltage value on the centerburst of the interferogram, which can be display in the computer screen as other parameters are optimized and sample alignment is performed (see the section 4c). It should be as high as possible.

- Iris opening (Aperture Setting): the iris opening is available with pre-selected sizes from 0.25 to 6 mm. It should be optimized to obtain maximum throughput while minimizing the beam size, to minimize the beam divergence.  The original IR beam is approximately 1" in diameter. Based on this value and the focal length of the focusing mirrors, a divergence range at the sample could be calculated.  Total light throughput should also be kept at values low enough so the signal is proportional to light intensity (there is a saturation of the detector at high light fluences).

- Scanning rate (Scanner Velocity): the most interesting values available are 10 or 20 kHz for DLaTGS detecter and 80 or 100 kHz for MCT detectors.  Faster scanning rates lead to faster data acquisition, but very fast scanning rates may lead to increases in S/N.  The maximum scanning rate should be chosen where noise levels are not increased (this can be evaluated by taking S/N measurements for the different scan rates using the same conditions).

- Number of scans (or Scan Times in minutes): to signal average.  S/N ratios should be increased as the square root of the number of scans, but too long data acquisition times lead to drifts in IR background and in changes in the nature of the sample (adsorption, etc.).

- Resolution: High resolutions are needed to separate different IR peaks, but require longer acquisition times (require further travel of the interferometer mirror), introduce noise (from the wings of the interferogram), and may reduce total peak signal intensity.  Typical value is 4 cm-1, sufficient for surface adsorbates, but sometimes 2 cm-1 is required.

- Amplification (Signal Gain): there are preamplifier gains (Automatic, (1, (2, (4, (8, or (16), to be set to optimize signal without saturating the centerburst.  The gain at the centerburst may be set separately (at a lower value) than for the rest (wings) of the interferogram.

## d.	Detector

### i.	Types

There are three different types of IR detectors: thermal, pyroelectric, and photoconducting detectors. Thermal detector such as thermocouple or bolometer is used over a wide range of wavelengths at room temperature, but not preferred due to slower response time and lower sensitivity than other types of detectors. Pyroelectric detector has a much faster response time as it depends on the rate of temperature change. The most common material is “deuterium L-alanine doped triglycene sulphate (DLaTGS). Photoconducting detector relies on the interaction between photons and a semiconductor such as mercury cadmium telluride (MCT) or indium antimonide (InSb), so response time is much faster and sensitivity is also much higher. Usually it has to be cooled first prior to use for thermal noise. Tensor 27 for DRIFT uses DLaTGS detector on the internal beam line.

| Detector | Range (cm-1) | Sensitivity (cm Hz1/2 W-1) | Cooling |
| --- | --- | --- | --- |
| DLaTGS | 12,000–350 | >4 x 108 | Room Temperature |
| MCT Wide Band | 12,000–420 | >5 x 109 | LN2 cooling |
| MCT Medium Band | 12,000–600 | >2 x 1010 | LN2 cooling |
| MCT Narrow Band | 12,000–850 | >4 x 1010 | LN2 cooling |
| InSb | 12,800–1850 | >1.5 x 1011 | LN2 cooling |

Figure 1. Detector compartment of Tensor 27 FTIR spectrometer.

### ii.	Substitution

1. 	Loosen the fixing screw of the detector compartment cover by using 6 mm Allen wrench.

- Open the detector compartment.

- Loosen the fixing screw of the detector with 6 mm Allen screw wrench (see Figure 1).

- Pull the detector straight upward out of the dovetail guide.

- Insert the detector you want into the dovetail guide and press it right down.

- Fasten the Allen screw.

- Close the detector compartment.

- Check the signal intensity using the OPUS program.

### iii.	Preparation

MCT or InSb detector has to be cooled down to liquid nitrogen temperature prior to use. Always wear blue Cryo-gloves (Tempshield) when you handle liquid nitrogen (please refer SOP of cryogenic materials in the “Lab Safety Manual” folder or download “SOPs Process” (http://research.chem.ucr.edu/groups/zaera/images/documents/4_sop_processes_2013-07.pdf) from our group homepage. Cool the detector with LN2 using the funnel fill tube. Keep filling the funnel in cone increments. Repeat until reservoir is full (spill over). Cool down time is about 5 to 10 minutes. The detector is at working temperature once the rapid boil off venting has stopped. Place the plug (cap) in the fill port

### iv.	Maintenance

The LN2 Dewar may require a vacuum refreshing. If the vacuum level of LN2 dewar will be decreased below acceptable, then the surface of that vacuum jacket is very cold. Sometimes it is covered with ice because of the condensation of moister form the air, so water peaks appear on the spectra.

            - Remove the MCT detector from the spectrometer.

            - Connect the adapter (Figure 2) to one of vacuum valves.

            - Switch on the vacuum pump and open the valve.

            - Inspect the O-ring inside the adapter (B in Figure 2) for signs of wear.

            - Remove the cap from the connection nozzle of the detector.

            - Pull the adapter knob (D in Figure 2).

            - Loosen the coupling nut (B in Figure 2).

            - Push the adapter carefully over the connection nozzle of detector dewar.

            - Fasten the coupling nut hand-tight while holding the adapter and detector.

            - Push the adapter knob in the closed position until the threaded rod of the adapter is in contact with the sealing plug of the dewar evacuation valve.

            - Evacuate the detector dewar.

            - Close the vacuum valve.

            - Screw the threaded rod of the adapter in the connection thread of dewar evacuation valve by turning the adapter knob clockwise.

            - Pull the knob to the open position in order to open the dewar evacuation valve.

            - Evacuate the detector by opening the vacuum valve to pressure less than 10-6 Torr.

            - Close the dewar evacuation valve by pushing the adapter knob (D in Figure 2).

            - Press adapter knob firmly to the stop position.

            - Screw the threaded rod of the adapter out of the connection thread of the dewar evacuation valve by rotating the adapter knob.

            - Vent the section between vacuum pump and adapter.

            - Pull the knob to the open position.

            - Loosen the coupling nut and remove the adapter from the connection nozzle of the detector dewar.

            - Reinstall the MCT detector in the spectrometer.

Figure 2. Adapter with flexible metal hose and Swagelok connection.

## e.	Optical Alignment

Tensor 27 is equipped with a high stability interferometer with ROCKSOLID permanent alignment. Also detector substitution is possible with Tensor 27. DLaTGS detector is default, and optional MCT detectors are available in our lab. After substituting the detectors, re-alignment is not necessary due to the dovetail detector mounting. However, it has to be done properly whenever any accessory like DRIFTS or ATR system is installed or uninstalled in the sample compartment. Adjustment knobs for the alignment are shown in Figure 1.

            - Remove all the accessory and samples from the IR beam path. Make sure that there is nothing in the sample compartment.

            - MCT detector has to be cooled first prior to measurement.

            - Run OPUS program.

            - Select the “Measurement” menu.

            - Click the “Check Signal” tab.

            - Select the “Interferogram” radio button.

            - The amplitude value indicates the signal intensity that is currently detected.

            - Loosen the fixing screw of the detector compartment cover by using 6 mm Allen wrench.

            - Open the detector compartment.

            - Adjust two knobs (2 and 3 in Figure 1) of the mirror to get the highest value while watching the amplitude on the monitor. If you lost IR beam, use IR card to find it.

            - Once you get a good alignment, close the cover.

            - Wait for 6 hrs at least to be purged completely.

## f. 	Transmittance spectra

### i.	Choice of Background Spectra

Depending on how to prepare your sample, it is strongly recommended to perform the background measurement with a proper material. The purpose of the background measurement is to detect the influence of the ambient conditions (e.g. humidity or temperature) and the auxiliary materials (e.g. solvents) that are required for preparing the sample. The subsequent sample measurement results in the sample spectrum from which the influence is eliminated. So, do both background and sample measurements with the same parameter settings in OPUS program. Ensure that the ambient conditions are identical or at least nearly identical.

## g.	OPUS Software

### i.	Spectra Acquisition

- Start the OPUS program:

  - Click the short-cut icon of the OPUS program on the desktop.

  - The OPUS login window will appear on the screen.

  - Type "OPUS" in the password line (for default operator).

- Select "Measure" and then "Measurement". After a few seconds, the measurement set-up windows will appear on the screen. Check the experimental parameters under "Advanced", "Optics", "Acquisition", "FT", and "Display".

- Set the sample temperature as needed.

- When the sample reaches the desired temperature, take a background spectrum by selecting "Measurement", "Basic", and "Collect Background". The program will display a message in the bottom line high-lighted by a green color indicating that the data are being taken.

- When the message disappears, the measurement is done.  From this point on, you can proceed with the experiments (e.g. heating of the sample, adsorption of gases etc.).

- During the data acquisition, the interferogram cannot be seen on the check signal menu. After the background scans are finished, the "Measurement" window remains on the screen.

- Go to “Background” and Click “Save Background” to save the spectrum. The Filename (####.0) is saved automatically on the hard drive and displayed on the left (Window List) of the screen.

- The file can be saved as other formats on the hard drive. Select "Save File As", then check the file name and path to be saved. If you want an ASCII file, in the "Mode" option on the Save Spectrum window, select "Data Point Table" before clicking the "Save" button.

- Fill the chamber with the desired pressure of the reactant gases, at the same temperature used for the acquisition of the background spectrum.

- Take the sample spectra by selecting "Collect Sample", and wait until the highlighted message disappears.

- To take the spectra with different backgrounds saved previously, select “load background” and click any background spectrum.

- Type the sample name and explanation in the lines for sample name and sample form. (e.g. Sample Name: 1% Pt/SiO2; Sample Form: after CO dosing of 10 Torr for 5 min at -150oC)

- Take the sample data by selecting "Collect Sample" and wait until the highlighted message disappears.

- Select “load background” again, and click another background spectrum for the next experiment.

Figure. 3. “Manipulate” and “Evaluate” menus in OPUS program.

### ii.	Data Processing

Very useful processing tools are available in OPUS program. For example, Base Line Correction and Spectrum Calculator are in “Manipulate” menu, while Curve Fit and Integration are in “Evaluate” menu.  Use helpful menu to learn how to do them.

## h. 	Maintenance

### i.	IR Source

Tensor 27 uses only one mid IR source (MIR) in an U-shaped silicon carbide piece, which is pre-aligned, so realignment is not required after replacement. During the spectrometer operation, it becomes very hot and has risk of skin burn. Please avoid any skin contact and wait until the source cools down enough before you remove it.

Beam Detection

If you like to check the IR beam visually, you can use an IR sensor card (see the right image) supplied from Bruker when the spectrometer was installed. You may confuse the red light on your sample with the IR beam, but actually it is He-Ne laser light (632.8 nm) that travels along the IR beam path.

Intensity Measurement

IR beam signal can be checked on the OPUS program.

1. 	Remove all the accessory and samples from the IR beam path. Make sure that there is nothing in the sample compartment.

2. 	Run OPUS program

3. 	Select the “Measurement” menu

4. 	Click the “Check Signal” tab

5. 	Select the “Interferogram” radio button

6. 	The amplitude value indicates the signal intensity that is currently detected.

Cooling

Tensor 27 uses air-cooled MIR source, so additional cooling system isn’t required for normal operation. However, please do not block the ventilation slot of the source/laser compartment at the spectrometer top-side. During the spectrometer operation, the IR source generates heat, which is dissipated by the slots. Failure to do so can lead to spectrometer component damage.

Equinox uses a water-cooled MIR source with a cooling system (Aquatherm Model # WLR 200), so its water level has to be checked frequently.

Figure 4. OPUS dialog window for “Measurement”.

Replacement

            - Switch off the spectrometer.

            - Wait until the source cools down sufficiently.

            - Open the source/laser compartment.

            - Loosen the Knurled thumb screw of the clamping bar (approximately one turn).

            - Press the source downward while rotating the clamping bar aside.

            - Take the source out of the holder.

            - Insert a new source in the operating position holder.

            - While pressing the source downward, rotate the clamping bar over the source.

            - Tighten the knurled thumb screw of the clamping bar about one turn.

            - Close the source/laser compartment.

            - Switch on the spectrometer.

            - Check the signal intensity using OPUS program.

            - Select “Optics Diagnostics” in Measure menu.

            - Click the source icon on the “Instrument Status” window.

            - Click on “Service Info” button.

            - Click on the “Reset” button (Figure 6).

Figure 5. Source/Laser compartment.

Figure 6. Source diagnostics page.

### ii.	HeNe Laser

FT-IR spectrometer is equipped with a He-Ne laser (632.8 nm), which controls the position of the moving mirror in interferometer for IR beam scanning. If you have a problem with the laser, a red STATUS indicator or a failed performance qualification test (PQ).

Figure 7. Source/Laser compartment.

- Measuring and Adjusting Signal Intensity

If you have laser problem, you need to check the laser amplitude. For this test you nee an oscilloscope. There are two test points (LA2 and LB2) on IFM board (G in Figure 7) You should see minimum of 2 V p-p sign wave on these points.

There are parameters that can and should be checked with respect to the HeNe laser intensity.

- HeNe laser output intensity – this value can be checked under measurement. Go “Optics Setup and Service Option”. Select “Service” tab, if the log window is blank. Select “Repeat Diagnostics Test” button, when the results appear take note of the current and initial laser intensity values as well as the number of laser dropout value.

- HeNe laser modulation intensity – this value can be obtained via the measurement. Go “Direct Command Entry Option”. In there type; VEL=6 <Enter>, SCM=3 <Enter>, LAA <Enter>. Do this three times and record the approximate average value given (the value is the peak to peak modulation amplitude in millivolts). LAB <Enter>. Do this three times and record the approximate average value given (the value is the peak to peak modulation amplitude in milllivolts).

Replacement

In case of defective laser you have to replace the complete laser module (C in Figure 7), which consists of the laser tube and the laser power supply.

            - Switch off the spectrometer.

            - Unplug the main power cable.

            - Open the source/laser compartment.

            - Loosen the Allen screw (A in Figure 7)

            - Rotate the holding plate (B in Figure 7)

            - Take the laser module out of the holder. Be aware of the fact that the laser module is still connected to the laser supply cable (D in Figure 7).

            - Loosen the two slotted screws (E in Figure 7) and unplug the supply cable (D in Figure 7).

            - Connect the supply cable to the replacement laser module.

            - Fasten the two slotted screws.

            - Open the front shutter of the laser tube if it is closed.

            - Insert the laser module in the holder. Make sure that the laser rests on the holder bottom in a plane manner to ensure the correct orientation of the laser beam.

            - Rotate the holding plate (B in Figure 7) in place.

            - Fasten the Allen screw (A in Figure 7).

            - Close the source/laser compartment.

            - Plug the main power cord

            - Switch on the spectrometer.

            - Check the laser indicator on the spectrometer topside lights yellow after a few seconds. Otherwise, solve this issue.

            - Wait until the STATUS indicator lights green. If the indicator remains red, the laser module is not installed correctly. Correct the laser installation.

            - Check whether the IR signal is detected using the OPUS program.

            - Run the “Optics Diagnostics” in OPUS

            - Open “Instrument Status” window.

            - Click “HeNe laser” icon.

            - Click “Service Infor” button.

            - Click “Reset” button in the laser diagnostics page (Figure 8).

            - Perform an OQ Test using OVP (OPUS Validation Program).

Figure 8. Laser diagnostics page.

### iii.	Gas Purging

At least, the optical bench has to be purged with dry air or nitrogen gas as the beamsplitter is made of KBr, which is a hygroscopic material. Also, purging the FTIR spectrometer reduces the content of undesired atmospheric interference from moisture and carbon dioxide, which absorb IR light in both detector and sample compartments as well as the optical bench.

Right now the DRIFTS system is purged with nitrogen gas. The flow rate is controlled by two rotameters (right images): one is for the optical bench and the detector compartment; the other is for sample compartment. One or two nitrogen tanks are consumed per week.

# 3.	Gas Handling System

Gas manifold of DRIFTS system includes Ar, H2, O2, and reactant gas lines. The pressure is monitored by two sets of Baratron (9 and 10 in Figure 9) and thermocouple (11 and 12 in Figure 9) gauges. One set (9 and 11) is for DRIFTS system, and the other set (10 and 12) is for gas manifold.

## a. 	Design

Figure 9. Gas manifold for the DRIFTS system.

            - main valve to pump the low temperature reaction chamber (CHC)

            - auxiliary valve to pump the CHC

            - main valve to pump gas manifold

            - valve to supply gas

            - H2 gas valve

            - O2 gas valve

            - CO gas valve

            - Ar gas valve

            - Baratron gauge 1 to monitor the CHC

            - Baratron gauge 2 to monitor the gas manifold

            - thermocouple gauge 1 to monitor the CHC

            - thermocouple gauge 2 to monitor the gas manifold

            - gas-out line from the CHC

            - gas-in line form the CHC

            - Tube for gas reservoir

## b. 	General Operation Procedure

For Sample Change

            - Close valves #1, #2, and #4.

            - Open valve #3 to pump the tube (15 in Figure 9).

            - Close valve #3, and then fill the tube with Ar gas as desired.

            - Fill the CHC with Ar gas by opening valve #4 up to 800 Torr.

            - Close valve #4.

            - Open valve #3 to pump the tube.

            - Disassemble the CHC and change the sample.

            - Assemble the CHC.

            - Open valve #1 to pump the CHC.

For Gas Supply

            - Close valves #2, #3, and #4.

            - Fill the tube (15 in Figure 9) with gases you want.

            - Close valve #1.

            - Fill the CHC with the gases by opening valve #4 slowly.

            - Close valve #4.

            - Open valve #3 to pump the tube.

## c. 	Gas and Liquid Sample Handling

Gas samples is delivered in a lecture bottle or Sure/PacTM steel cylinder (Aldrich, right image), so use a proper regulator and Swagelok connection to the gas manifold. Liquid sample is placed in a glass container, which is specially designed by modifying glass-stainless steel tube adapter that facilitates connection to the gas manifold via Swagelok union. Freeze-pump-thaw cycles are required to remove air from the liquid sample prior to use.

## d. 	Maintenance

DRIFTS, Nanoreactor, and Praxis systems are sharing the three gas cylinders (O2, Ar, and H2) in Chemical Science building room 137 as seen in Figure 10, so please don’t be confused with them. When you need to use any gas, please check first if someone is using the gas line. If so, you need to get permission to switch the valve to your side.

Figure 10. O2, Ar, and H2 gas cylinders for DRIFTS, Nanoreactor and Praxis chambers.

## e. 	Valves

Bellow-sealed BK series (e.g. SS-4BK) from Swagelok are preferred for the valves in the manifold system. The valves can be disassembled for cleaning inside, and the PCTFE stem tip is also replaceable.

## f. 	Pressure Gauges

Two different types of pressure gauges are used in the gas manifold. One is Baratron gauge with digital readout for the range from 0 to 1000 Torr to control the amount of reactant gases quantitatively. Channel 1 (CH 1) is to the reaction chamber side, and channel 2 (CH 2) is to the gas manifold side. The other is thermocouple gauge to monitor vacuum status in the gas manifold and the reaction chamber.

## g. 	Mechanical Pumps

A mechanical pump is used for low vacuum in the gas manifold. Please check the oil level, and refill when the level is too low or replace dirty oil when the color is too dark. The process for handling vacuum oil is available in the SOP of  “Pump Oil”. Refer it in our “Lab Safety Manual” or download “SOPs Process” (http://research.chem.ucr.edu/groups/zaera/images/documents/4_sop_processes_2013-07.pdf)

# 4.	Praying MentisTM DRIFTS Cell

## a.	General Description

The Praying Mentis system is specifically designed for examining powders by diffuse reflection spectroscopy. The optical design of the Praying MentisTM incorporates two 6(, 90o off-axis ellipsoid mirrors. The ellipsoids are optically arranged so that their near focal points coincide and they are arranged to discriminate against the collection of reflected radiation. This is achieved by placing the collecting ellipsoid 60° away from the direction of the reflected radiation. The sample is placed in the common focal point of the two ellipsoid mirrors. The beam is condensed to a spot less than 2 mm in diameter. This allows analysis of very small quantities of sample.

## b.	Installation

### i.	Open Beam Spectrum

Prior to installation, make sure all sample holders are removed from the sample compartment. Collect an open beam background spectrum without accessory in the sample compartment. This spectrum should be used later to verify the throughput of the Praying MantisTM

### ii.	Purge Installation

Install the Praying MentisTM onto the floor of the sample compartment using the base plate provided. Extended the purge sleeves to the walls of the spectrometer and tighten the thumbscrews on the purge collars.

## c.	Alignment

### i.	Alignment

- Open the door on the front of the attachment.

- Slide the alignment fixture into the Praying MantisTM, with the horizontal mirror going in first. In this orientation, the tilted mirror is in the sampling position.

- Gently push the alignment fixture in against the stop.

- Remove the plastic caps over mirror M2 and M5. This allows access to the turn and tilt adjustments for these mirrors.

- Set the spectrometer to measure the “energy” on the detector.

- Adjust the turn and tilt controls for M5 with the supplied 3/32” hex driver until the signal on the detector is maximized.

- Adjust the height of the alignment fixture with the height adjustment knob to maximize the signal.

- Adjust the turn and tilt controls for M2. Repeat this sequence until there is no further increase in the signal on the detector.

### ii.	Verifying Throughput

- Make sure that the specified background spectrum is the previously collected open beam background .

- Collect a transmittance spectrum with the Praying MentisTM in the sample.

- Read the maximum value at 2500 cm-1.

- The Praying MentisTM throughput at that wavenumber should be at least 40% and the spectrum should resemble the Praying MentisTM throughput spectrum shown below.

- Slide the alignment fixture into the Praying MentisTM, with the tilted mirror going in first. In this orientation, the horizontal mirror is in the sampling position.

- Gently push the alignment fixture in against the stop.

- The detector now measures the specular reflectance from the mirror. The energy should be near-zero. If it is not less than 10% of the throughput measured earlier, realign with the tilted mirror or contact Harrick Scientific.

## d.	Low-Temperature Reaction Chamber (CHC) Installation

### i.	Introduction

The Low Temperature Reaction Chamber (CHC) is designed for operation from -150oC up to 600oC under vacuum. The pressure range is from high vacuum (10-6 Torr) to pressures up to 103 Torr. At high pressures, the maximum operating temperature may be lower.

The CHC is made of chemically-resistant 316 stainless steel. Within the chamber is a temperature-controlled sample stage with an integral sample cup. This stage incorporates a cartridge heater and two thermocouples, one in the sampling stage and the other in the sample cup. The sample stage is thermally isolated from the outer chamber wall. A water-cooling jacket controls the temperature of the outer surface of the chamber and windows during high or low temperature operation. The chamber has three gas ports for evacuating, pressurizing or flowing gas through the sample. These ports are equipped with 1/4” VCO fittings. One of the ports leads directly under the sample cup, the other two lead into the sides of the chamber. The chamber is enclosed with a dome with three windows, two for the spectrometer radiation to enter and exit the chamber and third for viewing, illuminating, or irradiating the sample. This enables the use of the reaction chamber for photochemical studies. The standard material for the observation window is UV quartz and KBr for the other two windows.

### ii.	Observation Window Installation

The observation window is the window that is located between the two semicircular indentations on the rim of the dome.

- Insert an O-ring in one of the window ports.

- Hold the retaining ring upside-down and place a PTFE washer inside.

- Then carefully place a window on top of the PTFE washer.

- Hold the retaining ring so that it continues to retain the window and thread it into the window port.

- Tighten to secure the window in place using the tool provided for tightening and loosening the retaining rings.

- Repeat these steps until all three windows are installed.

- Be careful not to touch the windows during the installation.

### iii.	Dome Installation

- Prior to installation of the dome, make sure that the O-ring and groove on the bottom of the dome are clean and free of dust.

- Retract the retaining plates on the CHC as far as possible.

- Install the dome making sure that the orientation slot on the dome engages the orientation pin on the CHC.

- Slide the retaining plates all the way over the rim of the dome.

### iv.	CHC Installation

Before installing the CHC in the Praying MantisTM, the Praying MantisTM should have been aligned in the spectrometer.

- Unscrew the thumbscrew on the top plate of the Praying MantisTM.

- Carefully swing the top plate of the Praying MantisTM out of the way.

- Grasp the sample stage and unscrew its mounting screw.

- Slide the sample stage slightly to the left to disengage its mounting pins. .

- Remove the sample stage.

- Lift the accessory out of the sample compartment and rest it on its back.

- Loosen the two setscrews that secure the knob and slide the knob off its rod.

- Slide the locking adapter over the rod.

- Tighten the two setscrews to secure the locking adapter.

- Mount the CHC mounting plate on the Praying MantisTM body. Be sure to carefully slide the mounting plate between the bottom plate of the Praying MantisTM and the camlock plate.

- Slide the CHC mounting plate under the Praying MantisTM body locating the CHC mounting plate pin in the hole in the bottom of the Praying MantisTM.

- Use the two pan head screws provided to secure the mounting plate in place.

- Mount the CHC by guiding the locating pin, which is directly under CHC body into the locating hold in the base.

- Fasten the thumbscrew to secure the CHC to the CHC mounting plate.

### v.	Fluid / Electrical Connections

- If temperature control is desired, connect the thermocouple and the heater to the temperature controller (ATC).

- If the CHC is to be operated at temperatures above 100 oC or below 0 oC, use 1/4” tubing to connect the two cooling ports to a water inlet/outlet or a coolant circulator.

- If vacuum is desired, connect to the gas outlet (1/4” VCO fitting) to a vacuum pump using 3/8” I.D vacuum tubing.

- If reaction or process gases are to be passed through the sample, connect the gas inlet/outlet ports as needed to the gas sources. Connect the gas outlet and seal off any unused gas ports.

- If reaction or process gases are to be passed over the sample, seal off the gas outlet. Connect one of the gas inlet/outlet ports to the gas source and the other as the outlet.

## e.	Temperature Controller (ATC) Connection

### i.	Introduction

The Temperature Controller (ATC) can be manually programmed to reach a desired temperature or computer controlled for sophisticated temperature profiling. The controller automatically determines the best settings for optimal heating based on the thermal response of a particular system. It features two thermocouple inputs and one 24 V heater output, allowing regular or cascade operation. Regular or single loop control is used for accessories equipped with a sample thermocouple. Accessories equipped with a sample and heater thermocouple use cascade or double loop control. Cascade control minimizes damage to system components, allows for over sizing heaters for optimal heat-up rates and extends heater life by reducing thermal cycling of the heater.

### ii.	Connection

- Check to make sure that all the switches on the temperature controller are off.

- Plug in the AC line cord into the unit and into a suitable AC outlet.

- Connect the heater thermocouple to the connector labelled TC2 on the back of the temperature controller and the sample thermocouple to TC1.

- Connect the heater to the heater connector located on the front of the temperature controller.

### iii.	Setting Thermocouple Type

- Simultaneously press both the Up-arrow and the Down-arrow keys for approximately 3 seconds until the input menu prompt InPt appears in the upper display and the Setup Menu prompt SEt appears in the lower display: InPt / SEt

- Press the MODE key. In1 will appear in the lower display.

- If the upper display does not show an H, the symbol for the K-type thermocouple used in most Harrick accessories, press the Up-arrow key until it does. If the item to be heated does not use a K-type thermocouple, please see the Watlow manual for the correct setting.

- Press the MODE key five times. In2 will appear in the lower display.

- If the upper display does not show an H, the symbol for the K-type thermocouple used in most Harrick accessories, press the Up-arrow key until it does. If the item to be heated does not use a K-type thermocouple, please see the Watlow manual for the correct setting.

- Press the DSPY key to return to the main display format.

### iv.	Setting Control Type

- Simultaneously press both the Up-arrow and the Down-arrow keys for approximately 3 seconds until the input menu prompt InPt appears in the upper display and the Setup Menu prompt SEt appears in the lower display: InPt / SEt

- Press the Up-arrow key twice. 9LbL appears in the upper display: 9LbL / SEt

- Press the MODE key four times. Cntl will appear in the lower display.

- Press the Up-arrow key until CSCd appears in the upper display.

- Press the MODE key once more. CSAC will appear in the lower display.

- Press the Up-arrow key until dir appears in the upper display.

- Press the DSPY key to return to the main display format.

### v.	Setting Units

- Simultaneously press both the Up-arrow and the Down-arrow keys for approximately 3 seconds until the input menu prompt InPt appears in the upper display and the Setup Menu prompt SEt appears in the lower display: InPt / SEt

- Press the Up-arrow key twice. 9LbL appears in the upper display: 9LbL / SEt

- Press the MODE key once. C_F will appear in the lower display.

- Enter F for Fahrenheit units or C for Centigrade units.

- If the incorrect units are shown, press the Up-arrow key once to change to the temperature units.

- Press the DSPY key to return to the main display format.

### vi.	Setting the Low Temperature Limit

- Simultaneously press both the Up-arrow and the Down-arrow keys for approximately 3 seconds until the input menu prompt InPt appears in the upper display and the Setup Menu prompt SEt appears in the lower display: InPt / SEt

- Press the MODE key twice. rL1 will appear in the lower display.

- Use the Down-arrow key to set the upper display to -200.

- Press the MODE key five more times. rL2 will appear in the lower display.

- Use the Down-arrow key to set the upper display to -200.

- Press the DSPY key to return to the main display format.

### vii.	Setting the High Temperature Limit

- Simultaneously press both the Up-arrow and the Down-arrow keys for approximately 3 seconds until the input menu prompt InPt appears in the upper display and the Setup Menu prompt SEt appears in the lower display: InPt / SEt

- Press the MODE key three times. rH1 will appear in the lower display.

- Use the Up-arrow and/or Down-arrow keys to set the upper display to 750.

- Press the MODE key three times. rH2 will appear in the lower display.

- Use the Up-arrow and/or Down-arrow keys to set the upper display to 750.

- Press the DSPY key to return to the main display format.

### viii.	Setting the Low Power Limit

- Simultaneously press both the Up-arrow and the Down-arrow keys for approximately 3 seconds until the input menu prompt InPt appears in the upper display and the Setup Menu prompt SEt appears in the lower display: InPt / SEt

- Press the Up-arrow key twice. 9LbL appears in the upper display: 9LbL / SEt

- Press the MODE key nine times. LoP will appear in the lower display.

- Use the Down-arrow key to set the low power limit, shown in the upper display to 0.

- Press the DSPY key to return to the main display format.

### ix.	Setting the High Power Limit

- Simultaneously press both the Up-arrow and the Down-arrow keys for approximately 3 seconds until the input menu prompt InPt appears in the upper display and the Setup Menu prompt SEt appears in the lower display: InPt / SEt

- Press the Up-arrow key twice. 9LbL appears in the upper display: 9LbL / SEt

- Press the MODE key ten times. HiP will appear in the lower display.

- Use the Down-arrow key to set the low power limit, shown in the upper display to 35 or 75.

- Press the DSPY key to return to the main display format.

### x.	Setting the Auto-Tune Set Point

- Simultaneously press both the Up-arrow and the Down-arrow keys for approximately 3 seconds until the input menu prompt InPt appears in the upper display and the Setup Menu prompt SEt appears in the lower display: InPt / SEt

- Press the Up-arrow key twice. 9LbL appears in the upper display: 9LbL / SEt

- Press the MODE key eleven times. AtSP will appear in the lower display.

- Use the Up-arrow and/or the Down-arrow keys to set the set point percentage, shown in the upper display to 90.

- Press the DSPY key to return to the main display format.

### xi.	Auto-Tuning

- Use the Up-arrow and/or the Down-arrow keys to select the desired set point temperature. The set point is shown in the lower display. The actual temperature measured by the thermocouple is shown in the lower display.

- Make sure the accessory is at room temperature. The upper display should read between 18 and 26 degrees C.

- Press the MODE key twice. OFF will appear in the upper display and Aut will appear in the lower display.

- Press the Up-arrow key. PidA will appear in the upper display. Aut will alternate with tune in the lower display.

- Press display and turn the heater switch on. The set point will alternate with tune in the lower display and the actual temperature will appear in the upper display. The auto-tune process takes approximately 5 minutes. At the end of the auto tune process, the lower display shows the set point and the upper display shows the actual temperature, which should be maintained for an additional 5 minutes.

- Press the DSPY key to return to the main display format.

- When auto-tuning is complete, turn off the heater switch and allow the accessory to cool to room temperature. When the accessory is cool, the upper display should read between 18 and 26°C.

## f.	Sample Loading

### i.	Collecting Background Spectrum

- Place a sample screen in to the sample cup and locate the plastic overflow tray over the sample cup.

- Fill the sample cup with the reference material using the provided packing tool. Draw the straightedge across the top of the cup to level off the surface.

- Carefully remove the overflow tray.

- Install the dome.

- Swing the Praying MantisTM cover back into place and tighten the thumbscrew.

- Fill the Dewar with liquid nitrogen or other coolant.

- Set the temperature and pressure to desired operating conditions.

- Set the spectrometer to measure the “energy” on the detector.

- Adjust the elevation adjustment micrometer on the CHC to maximize the signal on the detector.

- Collect the background spectrum

### ii.	Sample Spectrum

- Restore the reference material to ambient conditions.

- Remove the dome.

- Empty the sample cup. This can be done by vacuuming out the reference sample or by removing the CHC from the Praying MantisTM and dumping out the material.

- If needed clean the sample cup.

- Fill the sample cup with the sample.

- Install the dome.

- Install the CHC into the Praying MantisTM if it was removed to empty the sample cup.

- Swing the top plate back into place and tighten the thumbscrew.

- Restore the desired sampling conditions.

- Collect the sample spectrum.

- Restore the sample to ambient conditions.

- Empty the sample cup and clean it before running the next sample.

## g.	Maintenance

### i.	Leak Checking

Each time the reaction chamber is reassembled for sample loading, leaks may develop due to improper setting of the O-rings and the pressure doesn’t reach below 30 mTorr within 6 hrs. If leaks occur, carefully inspect all O-rings and sealing surfaces for damage and contamination. Clean the O-ring and the surface and try to get vacuum again. Otherwise, use a helium leak detector to pinpoint small leaks.

### ii.	Replacing the O-ring

The Reaction Chamber have O-rings for sealing the windows, O-rings for sealing the access holes, O-rings on the fittings, and an O-ring to seal the dome to the Reaction Chamber. They can be replaced with new ones (ORV-023 or ORK-023) supplied from Harrick or McMaster-Carr.

### iii.	Cleaning the Reaction Chamber

- Vent the chamber to atmosphere.

- Remove the CHC from the Praying MantisTM, remove the dome, turn the CHC upside-down, and dump out sample trapped inside.

- Use a solvent to clean off the remaining sample from the holder and the thermocouple.

- Unscrew the four screws that secure the bottom cover of the CHC.

- Remove the bottom cover to reveal the access hole.

- Softly wipe or blow clean air through the hole to clean the interior of the CHC.

- When the interior is clean, make sure the O-ring on the bottom plate is clean and free of dirt.

- Re-install the bottom.

### iv.	Replacing the Heater

- Remove the CHC from the Praying MantisTM.

- Push a small, blunt object through the heater access hole at the opposite side of the heating wires.

- Slide the heater out the other side of the unit.

- To install a new heater, coat the heater with MgO powder (MaaloxTM or equivalent antacid).

- Slide it into the CHC until it hits the stop.

- Make sure that the heater is inserted completely into the body before re-securing the wires.

### v.	Replacing the Thermocouple

- Remove the CHC from the Praying MantisTM.

- Gently pull out the thermocouple.

- Slide the new thermocouple into the CHC until it hits the stop.

- Make sure that the thermocouple is inserted completely into the body before re-securing the wires.

# 5.	Typical Experiment Sequence

Below a description is provided of the procedure to generate a DRIFT spectrum of a powder sample.  Both a background spectrum, which is taken with the clean sample, and a sample spectrum, which is taken after adsorption (and sometimes reaction) of the molecules on the sample are needed. These background and sample scans are taken separately. The DRIFT data are collected on the PC using the Bruker OPUS software.

## a.	Initial Steps

- Sign-in in the logbook

- Vent CHC and load sample

- Reinstate vacuum environment

- Start up the OPUS program

## b.	Sample Preparation

- Clean the sample (if required) and start cooling down or heating up. Pay special attention to the cleanliness of the sample!  Since the sample data (to be taken later) will be divided by the background data that are obtained with the clean surface to generate the absorbance data (which are the data to be interpreted), any impurity or adsorbate that is present when the background data are taken will disturb the final absorbance spectrum.

- Start the OPUS program:

  - Click the short-cut icon of the OPUS program on the desktop.

  - The OPUS login window will appear on the screen.

  - Type "OPUS" in the password line (for default operator). For details on the program, see Section 2.e or consult the manual.

- Select "Measure" and then "Measurement". After a few seconds, the measurement set-up windows will appear on the screen. Check the experimental parameters under "Advanced", "Optics", "Acquisition", "FT", and "Display".

## c.	Background Spectrum Acquisition

- Cool down or heat up the sample to the desired temperature.

- When the sample reaches the desired temperature, take a background spectrum by selecting "Measurement", "Basic", and "Collect Background". The program will display a message in the bottom line high-lighted by a green color indicating that the data are being taken.

- When the message disappears, the measurement is done.  From this point on, you can proceed with the experiments (e.g. heating of the sample, adsorption of gases etc.).

- During the data acquisition, the interferogram cannot be seen on the check signal menu. After the background scans are finished, the "Measurement" window remains on the screen.

- Go to “Background” and Click “Save Background” to save the spectrum. The Filename (####.0) is saved automatically on the hard drive and displayed on the left (Window List) of the screen.

- The file can be saved as other formats on the hard drive. Select "Save File As", then check the file name and path to be saved. If you want an ASCII file, in the "Mode" option on the Save Spectrum window, select "Data Point Table" before clicking the "Save" button.

## d.	Gas Manifold Control

- Gas manifold is sharing argon, hydrogen, and oxygen gas cylinders with Nanoreactor and PRAXIS. Check them in use prior to turning the 3-way valve to DRIFT system.

- Fill the gas reservoir first with the first reactant gas while checking the pressure on Baratron gauge channel 1. It is safer and easier way to control lower pressures than atmosphere.

- Fill the Reaction Chamber as slowly as possible while monitoring Baratron gauge channel 2 because the powder in the sample cup can pop up as pressurized suddenly.

- Empty the reservoir for the next gas.

- Open the value slowly also to pump the chamber after measuring sample spectra.

## e.	Sample Spectrum Acquisition

- Fill the chamber with the desired pressure of the reactant gases, at the same temperature used for the acquisition of the background spectrum.

- Take the sample spectra by selecting "Collect Sample", and wait until the highlighted message disappears.

- To take the spectra, select “load background” and click any background spectrum.

- Type the sample name and explanation in the lines for sample name and sample form. (e.g. Sample Name: 1% Pt/SiO2; Sample Form: after CO dosing of 10 Torr for 5 min at -150oC)

- Take the sample data by selecting "Collect Sample" and wait until the highlighted message disappears.

- Select “load background” again, and click another background spectrum for the next experiment.

NOTE 1: Make sure that the sample position in the sample scan is exactly the same as in the background scan: Do not try to optimize the sample position between a background and sample scan, even if the signal intensity is lower than it was in the background scan!  Also, make sure that the temperature at which the sample is taken is the same as that at which the background was taken, so that the contraction of the manipulator is the same in both cases. Do not change any parameter between the background and sample scans.

NOTE 2: Keep accurate records of the data type in the files. Know exactly which files contain the sample ASCII data.

NOTE 3: Keep backups of the OPUS files only.  These files contain the actual data from which all other data can be easily derived.

## f.	Temperature Control

- To get the desired temperature, use arrow keys on the front panel of the ATC controller.

- Heating rate selection isn’t among the menus available in the ATC controller. Instead, selecting the high power limit properly is recommended for the temperature below 100oC. The lower the power limit is, the slower the heating rate will be.

- If the temperature fluctuates too much, lower the high power limit or do Auto-Tuning again.

## g.	Final Steps

- Clean the sample for the next experiment.

- Once finished with sample, Return to room temperature

- Vent CHC

- Unload sample

- Sign-out in the logbook

# 6.	Suggested Training for Beginner

For beginners to learn how to obtain consistent and reliable data, the following training is strongly recommended.  The data obtained from the suggested experiments should be contrasted with previous results for the same system, and checked by an experienced operator.

- 1 wt% Pt/SiO2 commercial catalyst is recommended for the standard sample.

- Use 60 mg of the catalyst or measure the amount you used.

- Obtain a background spectrum at room temperature after sample cleaning procedure.

- Fill the reaction chamber with 400 Torr of gas mixture (CO:O2=1:1)

- Measure sample spectra with 1 min interval up to 1 hr at the same temperature.

- Plot the CO2 peak area in the spectra as a function of time.

# 7.	Reference Materials & Contacts

## a.	Reference Materials

The Praying MantisTM User’s Manual (Manual Part No. DRP-M-04)

Temperature Controller (ATC-024-1) User’s Manual (Manual Part No. ATC-M-7)

Tensor User Manual, 7th updated edition, September 2011

## b. 	Contacts

Bruker Optics, Inc.

19 Fortune Dr. Billerica, MA 01821, USA

978-439-9899 (ext. 5227)

www.brukeroptics.com

info@brukeroptics.com

Harrick Scientific Product, Inc.

141 Tompkins Ave.

Pleasantville, NY 10570

800-248-3847, 9 am to 5 pm EST

www.harricksci.com

info@harricksci.com or techsupport@harricksci.com